QA/Test Report

Equipment name	:Spin coater 					Date:

Model No./ Serial no.	:

Customer/Job card No.	……………………………………………….

Standard Test Procedure:

(All the tests are performed based on the theory of the respective experiment and/or by calibrated instruments.)

| Sl No | Specific Test Conducted | Test Result | Specification/ Accepted values | Pass/Fail | Remarks |
| --- | --- | --- | --- | --- | --- |
| 1. | Basic Inspection |  | Model No., Painting , Finishing , stability and other visual parameters |  |  |
| 2. | Vacuum checking |  | Place the substrate and run at maximum RPM |  |  |
| 3. | Vibration and noise checking |  | Run at high rpm and observe noise /vibration |  |  |
| 4. | RPM testing using Tachometer |  | 1 % error maximum for all rpm |  |  |
|  |  |  | 60 rpm |  |  |
|  |  |  | 200 rpm |  |  |
|  |  |  | 1000 rpm |  |  |
|  |  |  | 2000 rpm |  |  |
|  |  |  | 3000 rpm |  |  |
|  |  |  | 4000 rpm |  |  |
|  |  |  | 5000 rpm |  |  |
|  |  |  | 7000 rpm |  |  |
|  |  |  | 9000 rpm |  |  |
|  |  |  | 10000 rpm |  |  |
| 5. | Acceleration testing By using a Tachometer |  | 10 rpm / sec -10 sec – 100 rpm |  |  |
|  |  |  | 50 rpm/sec-10 sec-500 rpm |  |  |
|  |  |  | 100 rpm/sec-10 sec-1000 rpm |  |  |
|  |  |  | 250 rpm/sec-10 sec-2500 rpm |  |  |
|  |  |  | 500 rpm/sec-10 sec-5000 rpm |  |  |
|  |  |  | 2000 rpm/sec-5 sec-10000 rpm |  |  |
| 5. | Keyboard/ Touch screen checking |  | Enter all the alpha numeric values and check its functioning |  |  |
| 6. | Test multi-step sample program |  | Verify the speed, acceleration and duration at each step. |  |  |
| 7. | Duration test |  | Run a sample program with below durations between the steps |  |  |
|  |  |  | 5 sec |  |  |
|  |  |  | 100 sec |  |  |
|  |  |  | 1000 sec |  |  |
| 7. | Lid switch testing |  | Machine stops working if we open lid while running (Also run a program by opening the lid) |  |  |
| 8. | UV led testing |  | By running a sample program check the UV led lighting |  |  |
| 8. | Accessories checking |  | AC cord, vacuum pump, Drain hose ,touch pen , user manual, software CD etc |  |  |

Checked by						Approved by

………………….							…………..	…………